## Slide 1: Using PICS for Monitoring Stability of Landsat 7 ETM+ and Landsat 8 OLI
19 July 2016
Nischal Mishra, Esad Micijevic, Md. Obaidul Haque
Stinger Ghaffarian Technologies (SGT), Contractor to the USGS EROS
esad.micijevic.ctr@usgs.gov, +1 (605) 594-2801

## Slide 2: Background
Landsat instruments’ responses are trended regularly over multiple PICS in Sahara desert
  Libya4, Libya1, Egypt1, Niger 1 & 2 and Sudan
Landsat 7 ETM+ temporal gains were updated in 2013 with 14 years of observations over PICS
Some anomalies were observed in relatively short trends in OLI data over several PICS, including the Libya 4 site
Analysis done to compare the trends observed in OLI data with ETM+ trends to find out how much PICS data are needed to make reliable estimates of drifts in sensor gains
Additional investigation using data from Landsat surface reflectance products

## Slide 3: PICS Data Processing
The PICS data are processed using Landsat Image Assessment System (IAS) PICS database
  Image pixel values (DNs) from ROIs prescribed by CNES are converted to TOA reflectance
  Per-pixel solar zenith angles are used in the TOA reflectance calculations
  Simple empirical BRDF model is used to account for the annual variations in the solar zenith angle
Linear regression is applied to the BRDF normalized TOA reflectance data to model the long term stability of the sensors
T-test is then used to determine the statistical significance of the estimated drift

## Slide 4: L7 ETM+ Trends

## Slide 5: Estimate of L7 ETM+ Gain Changes
Across different Sahara PICS, the observed drifts are well within 0.1% per year
The weighted average drift shows that the gain change is less than 0.05% per year!
Notes: updated

## Slide 6: Estimate of L8 OLI Gain Changes
PICS based estimates differ from on-board calibrators by ~0.5%
Libya 4 appears to be the biggest outlier!!

## Slide 7: PICS Trending: Libya 4
OLI
ETM+
ETM+ and OLI exhibit similar trends since 2013 (OLI launch)

## Slide 8: PICS Trending: Libya 4 (contd.)
OLI
ETM+

## Slide 9: Clouds Over Libya 4
16 Nov
18 Dec
03 Jan
19 Jan
20 Feb
7 Mar
23 Mar
8 Apr
10 May
More than 60% images acquired by L8 OLI over Libya 4 site since Nov 2015 were cloudy and not used in the analysis
Mostly true for other PICS, too

## Slide 10: Oil Fire Images from Libya (Sentinel-2)
Jan 6, 2016
Jan 7, 2016

## Slide 11: Aqua MODIS

## Slide 12: Observations
In most cases, stability of on-board calibrators is not within the uncertainties of PICS based measurements
  Offset between on-board calibrators and lunar calibration estimates vs. top PICS, such as Libya 4 and Sudan, is significant (~0.5% per year, over about 3.5 years!)
  The temporal uncertainties are still lowest in Libya 4 site, which may allude towards some systematic decrease in the responsivity of the site rather than the performance of the sensor

## Slide 13: Can L8 Surface Reflectance Products Answer Some of the Discrepancies?
Data products acquired from http://espa.cr.usgs.gov/
Stability estimates compared against standard Level1 products

## Slide 14: L8 OLI Surface Reflectance Trending
Observations
The stability estimates using surface reflectance products are very comparable to estimates from the standard Level1 products, for different test sites considered
However, these estimates are consistently ~0.5% off from the on-board calibrator measurements

## Slide 15: Stability of ETM+ After 2013
Since OLI launch, ETM+ shows similar behavior
  However, ETM+ is very stable, based on the lifetime statistics
  So, the gain changes observed in Libya 4 PICS are possibly due to the change in site characteristics rather than the sensor
Stability estimates of ETM+ over various 3-years time periods vary compared to lifetime stability
This would mean that time periods up to 3 years may not be adequate to average out all the noise sources over PICS including the site variability

## Slide 16: So, how much data do we need?
Analysis indicated that 3 years of data over PICS may not be enough to provide the reliable estimate of sensor stability
  Under the assumption that site is very stable
  Also, the level of change that we are trying to detect is really small, about 0.2% per year!
To answer the question, look at the data from ETM+ before any gain corrections were applied in 2013
    ETM+ gain change was calculated for every incremental 2 years since launch over Libya 4 and Sudan sites (i.e. 2 years, 4 years, 6 years and so on)

## Slide 17: ETM+ Drift Estimate for Different Time Periods
ETM+ gain was updated in 2013 with about 14 years of data over PICS
The uncertainty has dependence hence will decrease with time when the number of observations increase.
The trend indicates that the gain estimate started converging after 6-8 years for most bands

## Slide 18: ETM+ Gain Change Observed Over Sudan Site
Similar results with Sudan site
Gain trends started to converge in 6-8 years after launch, for most bands.

## Slide 19: Summary
PICS based stability estimates for OLI do not agree well with the on-board calibrators and the moon
  Differences of ~0.5% per year were observed
Even larger differences were noticed with Libya 4 site
  ETM+ showed similar discrepancy since 2013 which may indicate that the site is getting darker
  Analysis with ETM+ data indicated that about 6-8 years of data over PICS were needed for gain estimates to start converging
Some PICS need to be monitored carefully to identify possible long term and short term changes in the site characteristics